## GEX Form 100-269(a) Rev F
| Test Method: | 100-270
| Instrument S/N: |  |  | Photometric Performance at 465nm (values are in absorbance units)
| Instrument Model: | GENESYS 30 |  | Standard ID (Filter ID) | Certified Value (from cert.) | Uncertainty (from cert.) | G30 Spec (instrument spec) | Lower Limit (calculated) | Upper Limit (calculated) | Value As Found | Pass/Fail
| Date: | [10 empty] | 22
| Relative Humidty (%): |  |  | -1 |  |  | 0.002 | -0.002 | 0.002 |  | PASS
| Temperature (°C): |  |  | -2 |  |  | 0 | 0 | 0 |  | PASS
| Calibration Standards Set ID#: |  |  | -3 |  |  | 0 | 0 | 0 |  | PASS
| Calibration Certificate #: |  |  | -4 |  |  | 0 | 0 | 0 |  | PASS
Notes: |  |  |  | Photometric Performance at 590nm (values are in absorbance units)
|  |  |  | Standard ID (Filter ID) | Certified Value (from cert.) | Uncertainty (from cert.) | G30 Spec (instrument spec) | Lower Limit (calculated) | Upper Limit (calculated) | Value As Found | Pass/Fail
|  |  |  | -1 |  |  | 0.002 | -0.002 | 0.002 |  | PASS
|  |  |  | -2 |  |  | 0 | 0 | 0 |  | PASS
|  |  |  | -3 |  |  | 0 | 0 | 0 |  | PASS
|  |  |  | -4 |  |  | 0 | 0 | 0 |  | PASS
|  | Wavelength Accuracy Test at 525nm (nominal) |  |  |  |  | Wavelength Accuracy Test at 782nm (nominal)
[5 empty] | Measured Data (A) | (nm) |  |  |  | Measured Data (A) | (nm)
[6 empty] | -7 |  |  |  |  | -7
[6 empty] | -6 |  |  |  |  | -6
[6 empty] | -5 |  |  |  |  | -5
|  |  | Standard ID: | -5 |  | -4 |  | Standard ID: | -5 |  | -4
|  |  | Certified Value (nm): |  |  | -3 |  | Certified Value (nm): |  |  | -3
|  |  | Uncertainty (nm): |  |  | -2 |  | Uncertainty (nm): |  |  | -2
|  |  | G30 Spec (nm) | 2 |  | -1 |  | G30 Spec (nm) | 2 |  | -1
|  |  | Lower Limit (nm): | -2 |  | 0 |  | Lower Limit (nm): | -2 |  | 0
|  |  | Upper Limit (nm): | 2 |  | 1 |  | Upper Limit (nm): | 2 |  | 1
[6 empty] | 2 |  |  |  |  | 2
|  |  | Peak As Found (nm): | #N/A |  | 3 |  | Peak As Found (nm): | #N/A |  | 3
|  |  | Pass/Fail: | #N/A |  | 4 |  | Pass/Fail: | #N/A |  | 4
[6 empty] | 5 |  |  |  |  | 5
[6 empty] | 6 |  |  |  |  | 6
[6 empty] | 7 |  |  |  |  | 7
| Performed By: |  |  |  |  | Date:
ENTER NAME: Q.A. Technician
| Reviewed By: |  |  |  |  | Date:
ENTER NAME: Q.A. Manager

## GEX Form 100-269(b) Rev F
| Test Method: | 100-271 Evolution P.V. Method 1: Optical Standards and Hg Lamp
| Instrument S/N: |  | Photometric Performance at 465nm (values are in absorbance units)
| Instrument Model: |  | Standard ID (Filter ID) | Certified Value (from cert.) | Uncertainty (from cert.) | Evo Spec (instrument spec) | Lower Limit (calculated) | Upper Limit (calculated) | Value As Found | Pass/Fail
| Date: | [9 empty] | 22
| Relative Humidty (%): |  | -1 |  |  |  | 0 | 0 |  | PASS
| Temperature (°C): |  | -2 |  |  |  | 0 | 0 |  | PASS
| Calibration Standards Set ID#: |  | -3 |  |  |  | 0 | 0 |  | PASS
| Calibration Certificate #: |  | -4 |  |  |  | 0 | 0 |  | PASS
Notes: |  |  | Photometric Performance at 590nm (values are in absorbance units)
|  |  | Standard ID (Filter ID) | Certified Value (from cert.) | Uncertainty (from cert.) | Evo Spec (instrument spec) | Lower Limit (calculated) | Upper Limit (calculated) | Value As Found | Pass/Fail
|  |  | -1 |  |  |  | 0 | 0 |  | PASS
|  |  | -2 |  |  |  | 0 | 0 |  | PASS
|  |  | -3 |  |  |  | 0 | 0 |  | PASS
|  |  | -4 |  |  |  | 0 | 0 |  | PASS
| Performed By: |  |  |  | Date:
ENTER NAME - Q.A. Technician
| Reviewed By: |  |  |  | Date:
ENTER NAME - Q.A. Manager

## GEX Form 100-269(c) Rev F
| Test Method: | 100-271 Evolution P.V. Method 2 Optical Standards
| Instrument S/N: |  | Photometric Performance at 465nm (values are in absorbance units)
| Instrument Model: |  | Standard ID (Filter ID) | Certified Value (from cert.) | Uncertainty (from cert.) | Evo Spec (instrument spec) | Lower Limit (calculated) | Upper Limit (calculated) | Value As Found | Pass/Fail
| Date: | [9 empty] | 22
| Relative Humidty (%): |  | -1 |  |  |  | 0 | 0 |  | PASS
| Temperature (°C): |  | -2 |  |  |  | 0 | 0 |  | PASS
| Calibration Standards Set ID#: |  | -3 |  |  |  | 0 | 0 |  | PASS
| Calibration Certificate #: |  | -4 |  |  |  | 0 | 0 |  | PASS
Notes: |  |  | Photometric Performance at 590nm (values are in absorbance units)
|  |  | Standard ID (Filter ID) | Certified Value (from cert.) | Uncertainty (from cert.) | Evo Spec (instrument spec) | Lower Limit (calculated) | Upper Limit (calculated) | Value As Found | Pass/Fail
|  |  | -1 |  |  |  | 0 | 0 |  | PASS
|  |  | -2 |  |  |  | 0 | 0 |  | PASS
|  |  | -3 |  |  |  | 0 | 0 |  | PASS
|  |  | -4 |  |  |  | 0 | 0 |  | PASS
|  |  | Wavelength Accuracy Test at 525nm (nominal) |  |  |  | Wavelength Accuracy Test at 782nm (nominal)
|  |  |  | Standard ID: | -5 |  |  | Standard ID: | -5
|  |  |  | Certified Value (nm): |  |  |  | Certified Value (nm):
|  |  |  | Uncertainty (nm): |  |  |  | Uncertainty (nm):
|  |  |  | Evo220 Spec (nm) | 0.8 |  |  | Evo220 Spec (nm) | 0.8
|  |  |  | Lower Limit (nm): | -0.8 |  |  | Lower Limit (nm): | -0.8
|  |  |  | Upper Limit (nm): | 0.8 |  |  | Upper Limit (nm): | 0.8
|  |  |  | Valley As Found (nm): |  |  |  | Valley As Found (nm):
|  |  |  | Pass/Fail: | PASS |  |  | Pass/Fail: | PASS
| Performed By: |  |  |  | Date:
ENTER NAME - Q.A. Technician
| Reviewed By: |  |  |  | Date:
ENTER NAME - Q.A. Manager